## Abs Sensor Uncertainty
Design Values |  |  |  | Measurement Reconciliation |  | Value Used in Calculations |  |  | Available sensors
N1914A | Power Meter |  |  |  | Power Meter | N1914A | Power Meter | Rec | Model | Fmin MHz | Fmax GHz | Pmin dBm | Pmax dBm | Conn | Notes
E9300A-H25 | Sensor Model |  |  |  | Sensor Model | E9300A-H25 | Sensor Model | X | 8481A | 10 | 18 | -30 | 20 | N |  | Discontinued
1 | Frequency (GHz) |  |  |  | Frequency (GHz) | 1 | Frequency (GHz) |  | 8481B | 10 | 18 | 0 | 44 | N
21 | Power (dBm) |  |  |  | Power (dBm) | 21 | Power (dBm) | X | 8481D | 10 | 18 | -70 | -20 | N
1.4 | UUT VSWR (typ) |  |  |  | UUT VSWR (typ) | 1.4 | UUT VSWR (typ) |  | 8481H | 10 | 18 | -10 | 35 | N
Normal | Mode |  |  |  | Mode | Normal | Mode | X | 8482A | 0.1 | 4.2 | -30 | 20 | N |  | Discontinued
1 | Averages |  |  |  | Averages | 1 | Averages |  | 8482B | 0.1 | 4.2 | 0 | 44 | N
24 | Hours since zero |  |  |  | Hours since zero | 24 | Hours since zero |  | 8482H | 0.1 | 4.2 | -10 | 35 | N
[5 empty] | Cal Factor (ratio) | 1 | Cal Factor (ratio) | X | 8483A | 0.1 | 2 | -30 | 20 | N | 75 ohm
Based on Design Values | [8 empty] | 8483A-H03 | 0.1 | 3 | -30 | 20 | N | 75 ohm | (Cal factor probable uncertainty @ 3 GHz is 2.0%, sum of uncertainties is 3.7%)
1.13 | Sensor VSWR @ 50 MHz |  |  |  | Sensor VSWR @ 50 MHz | 1.13 | Sensor VSWR @ 50 MHz | X | 8485A | 50 | 26.5 | -30 | 20 | 3.5 mm |  | Discontinued
2.8 | Sensor Cal Factor Unc (%) |  |  |  | Sensor Cal Factor Unc (%) | 2.8 | Sensor Calibration Factor Unc (%) |  | 8485A-033 | 50 | 33 | -30 | 20 | 3.5 mm |  | Discontinued
1.13 | Sensor VSWR (spec) |  |  |  | Sensor VSWR (meas) | 1.13 | Sensor VSWR | X | 8485D | 50 | 26.5 | -70 | -20 | 3.5 mm
[9 empty] | 8485D-033 | 50 | 33 | -70 | -20 | 3.5 mm |  | Discontinued
Sanity Checks | [5 empty] | Power Meter |  | X | 8487A | 50 | 50 | -30 | 20 | 2.4 mm
Sensor | OK |  |  |  |  | 0.5% | Power Meter Instrumentation | X | 8487D | 50 | 50 | -70 | -20 | 2.4 mm
Power | OK |  |  |  |  | 0.4% | Calibrator Power Unc (%) |  | E4412A | 10 | 26.5 | -70 | 20 | N
Frequency | OK |  |  |  |  | 1.08 | Calibrator VSWR |  | E4413A | 50 | 26.5 | -70 | 20 | 3.5 mm
[6 empty] | 5.5 | Noise Multiplier |  | E9300A | 10 | 18 | -60 | 20 | N
| Absolute Power Sensor Measurement | [7 empty] | E9300A-H24 | 50 | 24 | -60 | 20 | 3.5 mm
| (Terms match power meter diagnostics) |  |  |  |  | Power Sensor |  | X | E9300A-H25 | 50 | 24 | -50 | 30 | 3.5 mm
Distribution | Term | u(x)/x (%) | u2(x)/x2 | u(x)/x (dB) | Source of uncertainty | E9300AH25Hi | HiLoSensor |  | E9300B | 10 | 18 | -30 | 44 | N |  | cal factors and return loss are characteristic > 18 GHz
Rayleigh | Mu | 0.2332% | 5.4361e-06 | 0.010114 | sensor/UUT mismatch | 0 | Additional Cal Factor Unc (%) |  | E9300H | 10 | 18 | -50 | 30 | N |  | cal factors and return loss are characteristic > 18 GHz
Rayleigh | Muc | 0.02807% | 7.8792e-08 | 0.0012189 | sensor/calibrator mismatch | 21 | Raw Power (dBm) |  | E9301A | 10 | 6 | -60 | 20 | N
Retangular | Pm | 0.2887% | 8.3333e-06 | 0.012519 | power meter instr (meas) | 0.12589 | Raw Power (W) |  | E9301B | 10 | 6 | -30 | 44 | N
Retangular | Pmc | 0.2887% | 8.3333e-06 | 0.012519 | power meter instr (cal) | 0.061033 | Sensor Rho |  | E9301B-H01 | 10 | 4.2 | -30 | 50 | N
Retangular | Pcal | 0.2309% | 5.3333e-06 | 0.010018 | calibrator output power | 0.061033 | Sensor Rho @ 50 MHz |  | E9301B-H50 | 10 | 6 | -30 | 47 | N
Retangular | D | 0.01651% | 2.7257e-08 | 0.00071695 | power meter drift | 0.038462 | Calibrator Rho |  | E9301H | 10 | 6 | -50 | 30 | N |  | VSWR 1.2 to 2 GHz, 1.3 to 3 GHz, 1.4 to 4.2 GHz
Retangular | Kb | 1.617% | 0.00026133 | 0.069646 | sensor calibration factor | 0.16667 | UUT Rho |  | E9304A | 0.009 | 6 | -60 | 20 | N
Retangular | Pl | 1.732% | 0.0003 | 0.074578 | power sensor linearity | 3% | Sensor Linearity (%) |  | E9304A-H01 | 0.009 | 18 | -30 | 44 | N |  | Specs 6 to 18 GHz undetermined, expect to be similar to E9300B
Retangular | Zsm | 0.002293% | 5.258e-10 | 9.9584e-05 | power meter zero set (meas) | 5e-06 | Sensor Noise (W, spec) |  | E9304A-H18 | 0.009 | 18 | -60 | 20 | N
Retangular | Zscal | 0.0002887% | 8.3333e-12 | 1.2537e-05 | power meter zero set (cal) | 5e-06 | Sensor Zero Set (W, spec) | X | E9304A-H19 | 0.009 | 18 | -50 | 30 | N |  | Specs 6 to 18 GHz undetermined, expect to be similar to E9300B
Gaussian | Nm | 0.01092% | 1.1929e-08 | 0.00047431 | sensor noise (meas) | 1.5e-06 | Sensor Zero Drift (W, spec) |  | E9304A-H20 | 0.009 | 18 | -60 | 20 | 3.5 mm |  | (Not sure what specs are - simply used E9304A-H18)
Gaussian | Ncal | 0.00035% | 1.225e-11 | 1.52e-05 | sensor noise (cal) | E9300AH25Lo | HiLoSensor - Calibration |  | N8481A | 10 | 18 | -35 | 20 | N
| Total | 2.427% | 0.00058889 | 0.10413 | Combined Uncertainty, k=1 | 7e-09 | Sensor Noise - Calibration |  | N8481A-CFT | 10 | 18 | -30 | 20 | N
|  | 4.853% |  | 0.20583 | Combined Uncertainty, k=2 | 5e-09 | Sensor Zero Set - Calibration |  | N8481B | 10 | 18 | -5 | 44 | N
[6 empty] | 0.001 | Calibration Power (W) |  | N8481B-CFT | 10 | 18 | 0 | 44 | N
[6 empty] | 0 | Calibration Power (dBm) |  | N8481H | 10 | 18 | -15 | 35 | N
[9 empty] | N8481H-CFT | 10 | 18 | -10 | 35 | N
[9 empty] | N8482A | 10 | 6 | -35 | 20 | N
Notes on mismatch uncertainty calculations | [8 empty] | N8482A-CFT | 10 | 6 | -30 | 20 | N
The general equation is sqrt(2)*rho1*rho2, where the values or rho are normalized | [8 empty] | N8482B | 0.1 | 6 | -5 | 44 | N
Sensor/UUT mismatch | [8 empty] | N8482B-CFT | 0.1 | 6 | 0 | 44 | N
The UUT rho is assumed to be a typical value, so is divided by 1/sqrt(2*ln(5)) | [8 empty] | N8482H | 0.1 | 6 | -15 | 35 | N
The sensor mismatch can be a hard spec (divided by 1/3.439) or measured (not adjusted) | [8 empty] | N8482H-CFT | 0.1 | 6 | -10 | 35 | N |  |  | Model has data in the sensor tables
Sensor/Calibrator mismatch | [8 empty] | N8485A | 10 | 26.5 | -35 | 20 | 3.5 mm |  |  | Model does not have data in the sensor tables
The calibrator rho is assumed to be a hard spec, so is divided by 3.439 | [8 empty] | N8485A-033 | 10 | 33 | -35 | 20 | 3.5 mm
The sensor mismatch can be a hard spec (divided by 1/3.439) or measured (not adjusted) | [8 empty] | N8485A-033CFT | 10 | 33 | -30 | 20 | 3.5 mm
[9 empty] | N8485A-CFT | 10 | 26.5 | -30 | 20 | 3.5 mm
[9 empty] | N8486AQ | 33000 | 50 | -35 | 20 | UG-383/U
[9 empty] | N8486AQ-CFT | 33000 | 50 | -30 | 20 | UG-383/U
[9 empty] | N8486AR | 26500 | 40 | -35 | 20 | UG-599/U
[9 empty] | N8486AR-CFT | 26500 | 40 | -30 | 20 | UG-599/U
[9 empty] | N8487A | 50 | 50 | -35 | 20 | 2.4 mm
[9 empty] | N8487A-CFT | 50 | 50 | -30 | 20 | 2.4 mm
[9 empty] | N8488A | 10 | 67 | -35 | 20 | 2.4 mm
[9 empty] | Q8486A | 33000 | 50 | -30 | 20 | UG-383/U
[9 empty] | Q8486D | 33000 | 50 | -70 | -20 | UG-383/U
[9 empty] | R8486A | 26500 | 40 | -30 | 20 | UG-599/U
[9 empty] | R8486D | 26500 | 40 | -70 | -20 | UG-599/U
[9 empty] | V8486A | 50000 | 75 | -30 | 20 | UG-385/U
[9 empty] | W8486A | 75000 | 110 | -30 | 20 | UG-387/U
Measurement Noise Multipliers
[9 empty] | Averages | Normal | x2
[9 empty] | 1 | 5.5 | 6.5
[9 empty] | 2 | 3.89 | 4.6
[9 empty] | 4 | 2.75 | 3.25
[9 empty] | 8 | 1.94 | 2.3
[9 empty] | 16 | 1 | 1.63
[9 empty] | 32 | 0.85 | 1
[9 empty] | 64 | 0.61 | 0.72
[9 empty] | 128 | 0.49 | 0.57
[9 empty] | 256 | 0.34 | 0.41
[9 empty] | 512 | 0.24 | 0.29
[9 empty] | 1024 | 0.17 | 0.2
Note on N2: Francis Small: Recommended lab standard
Note on F12: Francis Small: If provided, this value should equal sqrt(rho^2 + 2*u(rho)^2) to match the calculation in the Dobbert & Gorin paper.
Note on F14: Francis Small: If provided, this value should equal sqrt(rho^2 + 2*u(rho)^2) to match the calculation in the Dobbert & Gorin paper.

## Abs Sensor Tables
Sensor Constant Value Lookup Table
| EPM-P Series |  | EPM Series |  |  |  | P Series |  | EPM Series
| E4416A | E4417A | E4418A | E4418B | E4419A | E4419B | N1911A | N1912A | N1913A | N1914A
Instrumentation | 0.5% | 0.5% | 0.5% | 0.5% | 0.5% | 0.5% | 0.5% | 0.5% | 0.5% | 0.5%
Calibrator Accuracy | 0.6% | 0.6% | 0.6% | 0.6% | 0.6% | 0.6% | 0.4% | 0.4% | 0.4% | 0.4%
Calibrator SWR | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.08 | 1.08 | 1.08 | 1.08
Sensor Constant Value Lookup Table
| 8481A | 8481B | 8481D | 8481H | 8482A | 8482B | 8482H | 8483A | 8483A-H03 | 8485A | 8485A-033 | 8485D | 8485D-033 | 8487A | 8487D | E4412A | E4413A | E9300AH24Hi | E9300AH24Lo | E9300AH25Hi | E9300AH25Lo | E9300AHi | E9300ALo | E9300BHi | E9300BLo | E9300HHi | E9300HLo | E9301AHi | E9301ALo | E9301BH01Hi | E9301BH01Lo | E9301BH50Hi | E9301BH50Lo | E9301BHi | E9301BLo | E9301HHi | E9301HLo | E9304AH01Hi | E9304AH01Lo | E9304AH18Hi | E9304AH18Lo | E9304AH19Hi | E9304AH19Lo | E9304AH20Hi | E9304AH20Lo | E9304AHi | E9304ALo | N8481ACFT | N8481AHi | N8481ALo | N8481BCFT | N8481BHi | N8481BLo | N8481HCFT | N8481HHi | N8481HLo | N8482ACFT | N8482AHi | N8482ALo | N8482BCFT | N8482BHi | N8482BLo | N8482HCFT | N8482HHi | N8482HLo | N8485A033CFT | N8485A033Hi | N8485A033Lo | N8485ACFT | N8485AHi | N8485ALo | N8486AQCFT | N8486AQHi | N8486AQLo | N8486ARCFT | N8486ARHi | N8486ARLo | N8487ACFT | N8487AHi | N8487ALo | N8488AHi | N8488ALo | Q8486A | Q8486D | R8486A | R8486D | V8486A | W8486A
Zero Set | 5e-08 | 5e-05 | 2e-11 | 5e-06 | 5e-08 | 5e-05 | 5e-06 | 5e-08 | 5e-08 | 5e-08 | 5e-08 | 2e-11 | 2e-11 | 5e-08 | 2e-11 | 5e-11 | 5e-11 | 5e-07 | 5e-10 | 5e-06 | 5e-09 | 5e-07 | 5e-10 | 0.0005 | 5e-07 | 5e-06 | 5e-09 | 5e-07 | 5e-10 | 0.0005 | 5e-07 | 0.0005 | 5e-07 | 0.0005 | 5e-07 | 5e-06 | 5e-09 | 0.0005 | 5e-07 | 5e-07 | 5e-10 | 5e-06 | 5e-09 | 5e-07 | 5e-10 | 5e-07 | 5e-10 | 6.3e-08 | 6.3e-08 | 2.5e-08 | 6.3e-05 | 6.3e-05 | 2.5e-05 | 6.3e-06 | 6.3e-06 | 2.5e-06 | 6.3e-08 | 6.3e-08 | 2.5e-08 | 6.3e-05 | 6.3e-05 | 2.5e-05 | 6.3e-06 | 6.3e-06 | 2.5e-06 | 6.3e-08 | 6.3e-08 | 2.5e-08 | 6.3e-08 | 6.3e-08 | 2.5e-08 | 6.3e-08 | 6.3e-08 | 2.5e-08 | 6.3e-08 | 6.3e-08 | 2.5e-08 | 6.3e-08 | 6.3e-08 | 2.5e-08 | 6.3e-08 | 2.5e-08 | 5e-08 | 3e-11 | 5e-08 | 3e-11 | 2e-07 | 2e-07
Zero Drift | 1e-08 | 1e-05 | 4e-12 | 1e-06 | 1e-08 | 1e-05 | 1e-06 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 4e-12 | 4e-12 | 1e-08 | 4e-12 | 1.5e-11 | 1.5e-11 | 1.5e-07 | 1.5e-10 | 1.5e-06 | 1.5e-09 | 1.5e-07 | 1.5e-10 | 0.00015 | 1.5e-07 | 1.5e-06 | 1.5e-09 | 1.5e-07 | 1.5e-10 | 0.00015 | 1.5e-07 | 0.00015 | 1.5e-07 | 0.00015 | 1.5e-07 | 1.5e-06 | 1.5e-09 | 0.00015 | 1.5e-07 | 1.5e-07 | 1.5e-10 | 1.5e-06 | 1.5e-09 | 1.5e-07 | 1.5e-10 | 1.5e-07 | 1.5e-10 | 7e-09 | 7e-09 | 3e-09 | 7e-06 | 7e-06 | 3e-06 | 7e-07 | 7e-07 | 3e-07 | 7e-09 | 7e-09 | 3e-09 | 7e-06 | 7e-06 | 3e-06 | 7e-07 | 7e-07 | 3e-07 | 7e-09 | 7e-09 | 3e-09 | 7e-09 | 7e-09 | 3e-09 | 7e-09 | 7e-09 | 3e-09 | 7e-09 | 7e-09 | 3e-09 | 7e-09 | 7e-09 | 3e-09 | 7e-09 | 3e-09 | 1e-08 | 6e-12 | 1e-08 | 6e-12 | 4e-08 | 4e-08
Meas Noise | 1.1e-07 | 0.00011 | 4.5e-11 | 1e-05 | 1.1e-07 | 0.00011 | 1e-05 | 1.1e-07 | 1.1e-07 | 1.1e-07 | 1.1e-07 | 4.5e-11 | 4.5e-11 | 1.1e-07 | 4.5e-11 | 7e-11 | 7e-11 | 5e-07 | 7e-10 | 5e-06 | 7e-09 | 5e-07 | 7e-10 | 0.0005 | 7e-07 | 5e-06 | 7e-09 | 5e-07 | 7e-10 | 0.0005 | 7e-07 | 0.0005 | 7e-07 | 0.0005 | 7e-07 | 5e-06 | 7e-09 | 0.0005 | 7e-07 | 5e-07 | 7e-10 | 5e-06 | 7e-09 | 5e-07 | 7e-10 | 5e-07 | 7e-10 | 1.14e-07 | 1.14e-07 | 8e-08 | 0.000114 | 0.000114 | 8e-05 | 1.14e-05 | 1.14e-05 | 8e-06 | 1.14e-07 | 1.14e-07 | 8e-08 | 0.000114 | 0.000114 | 8e-05 | 1.14e-05 | 1.14e-05 | 8e-06 | 1.14e-07 | 1.14e-07 | 8e-08 | 1.14e-07 | 1.14e-07 | 8e-08 | 1.14e-07 | 1.14e-07 | 8e-08 | 1.14e-07 | 1.14e-07 | 8e-08 | 1.14e-07 | 1.14e-07 | 8e-08 | 1.14e-07 | 8e-08 | 1.1e-07 | 6.5e-11 | 1.1e-07 | 6.5e-11 | 4.5e-07 | 4.5e-07
Cal Power (W) | 0.001 | 0.001 | 1e-06 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 1e-06 | 1e-06 | 0.001 | 1e-06 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001
VSWR power ref | 1.06 | 1.06 | 1.05 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.05 | 1.05 | 1.06 | 1.05 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.06 | 1.05 | 1.06 | 1.05 | 1.06 | 1.06
Adtl Cal Fac Unc | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
Hi/Lo Sensor
Sensor
Power | 8481A | 8481B | 8481D | 8481H | 8482A | 8482B | 8482H | 8483A | 8483A-H03 | 8485A | 8485A-033 | 8485D | 8485D-033 | 8487A | 8487D | E4412A | E4413A | E9300A | E9300A-H24 | E9300A-H25 | E9300B | E9300H | E9301A | E9301B | E9301B-H01 | E9301B-H50 | E9301H | E9304A | E9304A-H01 | E9304A-H18 | E9304A-H19 | E9304A-H20 | N8481A | N8481A-CFT | N8481B | N8481B-CFT | N8481H | N8481H-CFT | N8482A | N8482A-CFT | N8482B | N8482B-CFT | N8482H | N8482H-CFT | N8485A | N8485A-033 | N8485A-033CFT | N8485A-CFT | N8486AQ | N8486AQ-CFT | N8486AR | N8486AR-CFT | N8487A | N8487A-CFT | N8488A | Q8486A | Q8486D | R8486A | R8486D | V8486A | W8486A
-999 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
-70 | OutOfRng | OutOfRng | 8481D | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 8485D | 8485D-033 | OutOfRng | 8487D | E4412A | E4413A | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | Q8486D | OutOfRng | R8486D | OutOfRng | OutOfRng
-60 | OutOfRng | OutOfRng | 8481D | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 8485D | 8485D-033 | OutOfRng | 8487D | E4412A | E4413A | E9300ALo | E9300AH24Lo | OutOfRng | OutOfRng | OutOfRng | E9301ALo | OutOfRng | OutOfRng | OutOfRng | OutOfRng | E9304ALo | OutOfRng | E9304AH18Lo | OutOfRng | E9304AH20Lo | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | Q8486D | OutOfRng | R8486D | OutOfRng | OutOfRng
-50 | OutOfRng | OutOfRng | 8481D | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 8485D | 8485D-033 | OutOfRng | 8487D | E4412A | E4413A | E9300ALo | E9300AH24Lo | E9300AH25Lo | OutOfRng | E9300HLo | E9301ALo | OutOfRng | OutOfRng | OutOfRng | E9301HLo | E9304ALo | OutOfRng | E9304AH18Lo | E9304AH19Lo | E9304AH20Lo | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | Q8486D | OutOfRng | R8486D | OutOfRng | OutOfRng
-35 | OutOfRng | OutOfRng | 8481D | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 8485D | 8485D-033 | OutOfRng | 8487D | E4412A | E4413A | E9300ALo | E9300AH24Lo | E9300AH25Lo | OutOfRng | E9300HLo | E9301ALo | OutOfRng | OutOfRng | OutOfRng | E9301HLo | E9304ALo | OutOfRng | E9304AH18Lo | E9304AH19Lo | E9304AH20Lo | N8481ALo | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | N8482ALo | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | N8485ALo | N8485A033Lo | OutOfRng | OutOfRng | N8486AQLo | OutOfRng | N8486ARLo | OutOfRng | N8487ALo | OutOfRng | N8488ALo | OutOfRng | Q8486D | OutOfRng | R8486D | OutOfRng | OutOfRng
-30 | 8481A | OutOfRng | 8481D | OutOfRng | 8482A | OutOfRng | OutOfRng | 8483A | 8483A-H03 | 8485A | 8485A-033 | 8485D | 8485D-033 | 8487A | 8487D | E4412A | E4413A | E9300ALo | E9300AH24Lo | E9300AH25Lo | E9300BLo | E9300HLo | E9301ALo | E9301BLo | E9301BH01Lo | E9301BLo | E9301HLo | E9304ALo | E9304AH01Lo | E9304AH18Lo | E9304AH19Lo | E9304AH20Lo | N8481ALo | N8481ACFT | OutOfRng | OutOfRng | OutOfRng | OutOfRng | N8482ALo | N8482ACFT | OutOfRng | OutOfRng | OutOfRng | OutOfRng | N8485ALo | N8485A033Lo | N8485A033CFT | N8485ACFT | N8486AQLo | N8486AQCFT | N8486ARLo | N8486ARCFT | N8487ALo | N8487ACFT | N8488ALo | Q8486A | Q8486D | R8486A | R8486D | V8486A | W8486A
-19.99 | 8481A | OutOfRng | OutOfRng | OutOfRng | 8482A | OutOfRng | OutOfRng | 8483A | 8483A-H03 | 8485A | 8485A-033 | OutOfRng | OutOfRng | 8487A | OutOfRng | E4412A | E4413A | E9300ALo | E9300AH24Lo | E9300AH25Lo | E9300BLo | E9300HLo | E9301ALo | E9301BLo | E9301BH01Lo | E9301BLo | E9301HLo | E9304ALo | E9304AH01Lo | E9304AH18Lo | E9304AH19Lo | E9304AH20Lo | N8481ALo | N8481ACFT | OutOfRng | OutOfRng | OutOfRng | OutOfRng | N8482ALo | N8482ACFT | OutOfRng | OutOfRng | OutOfRng | OutOfRng | N8485ALo | N8485A033Lo | N8485A033CFT | N8485ACFT | N8486AQLo | N8486AQCFT | N8486ARLo | N8486ARCFT | N8487ALo | N8487ACFT | N8488ALo | Q8486A | OutOfRng | R8486A | OutOfRng | V8486A | W8486A
-15 | 8481A | OutOfRng | OutOfRng | OutOfRng | 8482A | OutOfRng | OutOfRng | 8483A | 8483A-H03 | 8485A | 8485A-033 | OutOfRng | OutOfRng | 8487A | OutOfRng | E4412A | E4413A | E9300ALo | E9300AH24Lo | E9300AH25Lo | E9300BLo | E9300HLo | E9301ALo | E9301BLo | E9301BH01Lo | E9301BLo | E9301HLo | E9304ALo | E9304AH01Lo | E9304AH18Lo | E9304AH19Lo | E9304AH20Lo | N8481ALo | N8481ACFT | OutOfRng | OutOfRng | N8481HLo | OutOfRng | N8482ALo | N8482ACFT | OutOfRng | OutOfRng | N8482HLo | OutOfRng | N8485ALo | N8485A033Lo | N8485A033CFT | N8485ACFT | N8486AQLo | N8486AQCFT | N8486ARLo | N8486ARCFT | N8487ALo | N8487ACFT | N8488ALo | Q8486A | OutOfRng | R8486A | OutOfRng | V8486A | W8486A
-10 | 8481A | OutOfRng | OutOfRng | 8481H | 8482A | OutOfRng | 8482H | 8483A | 8483A-H03 | 8485A | 8485A-033 | OutOfRng | OutOfRng | 8487A | OutOfRng | E4412A | E4413A | E9300ALo | E9300AH24Lo | E9300AH25Lo | E9300BLo | E9300HLo | E9301ALo | E9301BLo | E9301BH01Lo | E9301BLo | E9301HLo | E9304ALo | E9304AH01Lo | E9304AH18Lo | E9304AH19Lo | E9304AH20Lo | N8481ALo | N8481ACFT | OutOfRng | OutOfRng | N8481HLo | N8481HCFT | N8482ALo | N8482ACFT | OutOfRng | OutOfRng | N8482HLo | N8482HCFT | N8485ALo | N8485A033Lo | N8485A033CFT | N8485ACFT | N8486AQLo | N8486AQCFT | N8486ARLo | N8486ARCFT | N8487ALo | N8487ACFT | N8488ALo | Q8486A | OutOfRng | R8486A | OutOfRng | V8486A | W8486A
-9.99 | 8481A | OutOfRng | OutOfRng | 8481H | 8482A | OutOfRng | 8482H | 8483A | 8483A-H03 | 8485A | 8485A-033 | OutOfRng | OutOfRng | 8487A | OutOfRng | E4412A | E4413A | E9300AHi | E9300AH24Hi | E9300AH25Lo | E9300BLo | E9300HLo | E9301AHi | E9301BLo | E9301BH01Lo | E9301BLo | E9301HLo | E9304AHi | E9304AH01Lo | E9304AH18Hi | E9304AH19Lo | E9304AH20Hi | N8481ALo | N8481ACFT | OutOfRng | OutOfRng | N8481HLo | N8481HCFT | N8482ALo | N8482ACFT | OutOfRng | OutOfRng | N8482HLo | N8482HCFT | N8485ALo | N8485A033Lo | N8485A033CFT | N8485ACFT | N8486AQLo | N8486AQCFT | N8486ARLo | N8486ARCFT | N8487ALo | N8487ACFT | N8488ALo | Q8486A | OutOfRng | R8486A | OutOfRng | V8486A | W8486A
-5 | 8481A | OutOfRng | OutOfRng | 8481H | 8482A | OutOfRng | 8482H | 8483A | 8483A-H03 | 8485A | 8485A-033 | OutOfRng | OutOfRng | 8487A | OutOfRng | E4412A | E4413A | E9300AHi | E9300AH24Hi | E9300AH25Lo | E9300BLo | E9300HLo | E9301AHi | E9301BLo | E9301BH01Lo | E9301BLo | E9301HLo | E9304AHi | E9304AH01Lo | E9304AH18Hi | E9304AH19Lo | E9304AH20Hi | N8481ALo | N8481ACFT | N8481BLo | OutOfRng | N8481HLo | N8481HCFT | N8482ALo | N8482ACFT | N8482BLo | OutOfRng | N8482HLo | N8482HCFT | N8485ALo | N8485A033Lo | N8485A033CFT | N8485ACFT | N8486AQLo | N8486AQCFT | N8486ARLo | N8486ARCFT | N8487ALo | N8487ACFT | N8488ALo | Q8486A | OutOfRng | R8486A | OutOfRng | V8486A | W8486A
-0.99 | 8481A | OutOfRng | OutOfRng | 8481H | 8482A | OutOfRng | 8482H | 8483A | 8483A-H03 | 8485A | 8485A-033 | OutOfRng | OutOfRng | 8487A | OutOfRng | E4412A | E4413A | E9300AHi | E9300AH24Hi | E9300AH25Lo | E9300BLo | E9300HLo | E9301AHi | E9301BLo | E9301BH01Lo | E9301BLo | E9301HLo | E9304AHi | E9304AH01Lo | E9304AH18Hi | E9304AH19Lo | E9304AH20Hi | N8481AHi | N8481ACFT | N8481BLo | OutOfRng | N8481HLo | N8481HCFT | N8482AHi | N8482ACFT | N8482BLo | OutOfRng | N8482HLo | N8482HCFT | N8485AHi | N8485A033Hi | N8485A033CFT | N8485ACFT | N8486AQHi | N8486AQCFT | N8486ARHi | N8486ARCFT | N8487AHi | N8487ACFT | N8488AHi | Q8486A | OutOfRng | R8486A | OutOfRng | V8486A | W8486A
0 | 8481A | 8481B | OutOfRng | 8481H | 8482A | 8482B | 8482H | 8483A | 8483A-H03 | 8485A | 8485A-033 | OutOfRng | OutOfRng | 8487A | OutOfRng | E4412A | E4413A | E9300AHi | E9300AH24Hi | E9300AH25Lo | E9300BLo | E9300HLo | E9301AHi | E9301BLo | E9301BH01Lo | E9301BLo | E9301HLo | E9304AHi | E9304AH01Lo | E9304AH18Hi | E9304AH19Lo | E9304AH20Hi | N8481AHi | N8481ACFT | N8481BLo | N8481BCFT | N8481HLo | N8481HCFT | N8482AHi | N8482ACFT | N8482BLo | N8482BCFT | N8482HLo | N8482HCFT | N8485AHi | N8485A033Hi | N8485A033CFT | N8485ACFT | N8486AQHi | N8486AQCFT | N8486ARHi | N8486ARCFT | N8487AHi | N8487ACFT | N8488AHi | Q8486A | OutOfRng | R8486A | OutOfRng | V8486A | W8486A
0.01 | 8481A | 8481B | OutOfRng | 8481H | 8482A | 8482B | 8482H | 8483A | 8483A-H03 | 8485A | 8485A-033 | OutOfRng | OutOfRng | 8487A | OutOfRng | E4412A | E4413A | E9300AHi | E9300AH24Hi | E9300AH25Hi | E9300BLo | E9300HHi | E9301AHi | E9301BLo | E9301BH01Lo | E9301BLo | E9301HHi | E9304AHi | E9304AH01Lo | E9304AH18Hi | E9304AH19Hi | E9304AH20Hi | N8481AHi | N8481ACFT | N8481BLo | N8481BCFT | N8481HLo | N8481HCFT | N8482AHi | N8482ACFT | N8482BLo | N8482BCFT | N8482HLo | N8482HCFT | N8485AHi | N8485A033Hi | N8485A033CFT | N8485ACFT | N8486AQHi | N8486AQCFT | N8486ARHi | N8486ARCFT | N8487AHi | N8487ACFT | N8488AHi | Q8486A | OutOfRng | R8486A | OutOfRng | V8486A | W8486A
17.01 | 8481A | 8481B | OutOfRng | 8481H | 8482A | 8482B | 8482H | 8483A | 8483A-H03 | 8485A | 8485A-033 | OutOfRng | OutOfRng | 8487A | OutOfRng | E4412A | E4413A | E9300AHi | E9300AH24Hi | E9300AH25Hi | E9300BLo | E9300HHi | E9301AHi | E9301BLo | E9301BH01Lo | E9301BLo | E9301HHi | E9304AHi | E9304AH01Lo | E9304AH18Hi | E9304AH19Hi | E9304AH20Hi | N8481AHi | N8481ACFT | N8481BLo | N8481BCFT | N8481HHi | N8481HCFT | N8482AHi | N8482ACFT | N8482BLo | N8482BCFT | N8482HHi | N8482HCFT | N8485AHi | N8485A033Hi | N8485A033CFT | N8485ACFT | N8486AQHi | N8486AQCFT | N8486ARHi | N8486ARCFT | N8487AHi | N8487ACFT | N8488AHi | Q8486A | OutOfRng | R8486A | OutOfRng | V8486A | W8486A
20 | 8481A | 8481B | OutOfRng | 8481H | 8482A | 8482B | 8482H | 8483A | 8483A-H03 | 8485A | 8485A-033 | OutOfRng | OutOfRng | 8487A | OutOfRng | E4412A | E4413A | E9300AHi | E9300AH24Hi | E9300AH25Hi | E9300BLo | E9300HHi | E9301AHi | E9301BLo | E9301BH01Lo | E9301BLo | E9301HHi | E9304AHi | E9304AH01Lo | E9304AH18Hi | E9304AH19Hi | E9304AH20Hi | N8481AHi | N8481ACFT | N8481BLo | N8481BCFT | N8481HHi | N8481HCFT | N8482AHi | N8482ACFT | N8482BLo | N8482BCFT | N8482HHi | N8482HCFT | N8485AHi | N8485A033Hi | N8485A033CFT | N8485ACFT | N8486AQHi | N8486AQCFT | N8486ARHi | N8486ARCFT | N8487AHi | N8487ACFT | N8488AHi | Q8486A | OutOfRng | R8486A | OutOfRng | V8486A | W8486A
20.01 | OutOfRng | 8481B | OutOfRng | 8481H | OutOfRng | 8482B | 8482H | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | E9300AH25Hi | E9300BHi | E9300HHi | OutOfRng | E9301BHi | E9301BH01Hi | E9301BHi | E9301HHi | OutOfRng | E9304AH01Hi | OutOfRng | E9304AH19Hi | OutOfRng | OutOfRng | OutOfRng | N8481BLo | N8481BCFT | N8481HHi | N8481HCFT | OutOfRng | OutOfRng | N8482BLo | N8482BCFT | N8482HHi | N8482HCFT | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
29.01 | OutOfRng | 8481B | OutOfRng | 8481H | OutOfRng | 8482B | 8482H | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | E9300AH25Hi | E9300BHi | E9300HHi | OutOfRng | E9301BHi | E9301BH01Hi | E9301BHi | E9301HHi | OutOfRng | E9304AH01Hi | OutOfRng | E9304AH19Hi | OutOfRng | OutOfRng | OutOfRng | N8481BHi | N8481BCFT | N8481HHi | N8481HCFT | OutOfRng | OutOfRng | N8482BHi | N8482BCFT | N8482HHi | N8482HCFT | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
30.01 | OutOfRng | 8481B | OutOfRng | 8481H | OutOfRng | 8482B | 8482H | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | E9300BHi | OutOfRng | OutOfRng | E9301BHi | E9301BH01Hi | E9301BHi | OutOfRng | OutOfRng | E9304AH01Hi | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | N8481BHi | N8481BCFT | N8481HHi | N8481HCFT | OutOfRng | OutOfRng | N8482BHi | N8482BCFT | N8482HHi | N8482HCFT | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
35.01 | OutOfRng | 8481B | OutOfRng | OutOfRng | OutOfRng | 8482B | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | E9300BHi | OutOfRng | OutOfRng | E9301BHi | E9301BH01Hi | E9301BHi | OutOfRng | OutOfRng | E9304AH01Hi | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | N8481BHi | N8481BCFT | OutOfRng | OutOfRng | OutOfRng | OutOfRng | N8482BHi | N8482BCFT | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
44.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | E9301BH01Hi | E9301BHi | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
47.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | E9301BH01Hi | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
50.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
Linearity (%)
Sensor
Power | 8481A | 8481B | 8481D | 8481H | 8482A | 8482B | 8482H | 8483A | 8483A-H03 | 8485A | 8485A-033 | 8485D | 8485D-033 | 8487A | 8487D | E4412A | E4413A | E9300A | E9300A-H24 | E9300A-H25 | E9300B | E9300H | E9301A | E9301B | E9301B-H01 | E9301B-H50 | E9301H | E9304A | E9304A-H01 | E9304A-H18 | E9304A-H19 | E9304A-H20 | N8481A | N8481A-CFT | N8481B | N8481B-CFT | N8481H | N8481H-CFT | N8482A | N8482A-CFT | N8482B | N8482B-CFT | N8482H | N8482H-CFT | N8485A | N8485A-033 | N8485A-033CFT | N8485A-CFT | N8486AQ | N8486AQ-CFT | N8486AR | N8486AR-CFT | N8487A | N8487A-CFT | N8488A | Q8486A | Q8486D | R8486A | R8486D | V8486A | W8486A
-999 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
-70 | OutOfRng | OutOfRng | 0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | 0 | OutOfRng | 0 | 3 | 3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | OutOfRng | 0 | OutOfRng | OutOfRng
-60 | OutOfRng | OutOfRng | 0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | 0 | OutOfRng | 0 | 3 | 3 | 3 | 3 | OutOfRng | OutOfRng | OutOfRng | 3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3 | OutOfRng | 3 | OutOfRng | 3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | OutOfRng | 0 | OutOfRng | OutOfRng
-50 | OutOfRng | OutOfRng | 0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | 0 | OutOfRng | 0 | 3 | 3 | 3 | 3 | 4 | OutOfRng | 4 | 3 | OutOfRng | OutOfRng | OutOfRng | 4 | 3 | OutOfRng | 3 | 4 | 3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | OutOfRng | 0 | OutOfRng | OutOfRng
-35 | OutOfRng | OutOfRng | 0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | 0 | OutOfRng | 0 | 3 | 3 | 3 | 3 | 4 | OutOfRng | 4 | 3 | OutOfRng | OutOfRng | OutOfRng | 4 | 3 | OutOfRng | 3 | 4 | 3 | 0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | 0 | OutOfRng | OutOfRng | 0 | OutOfRng | 0 | OutOfRng | 0 | OutOfRng | 0 | OutOfRng | 0 | OutOfRng | 0 | OutOfRng | OutOfRng
-30 | 0 | OutOfRng | 1 | OutOfRng | 0 | OutOfRng | OutOfRng | 0 | 0 | 0 | 0 | 2 | 2 | 0 | 2 | 3 | 3 | 3 | 3 | 4 | 3.5 | 4 | 3 | 3.5 | 3.5 | 3.5 | 4 | 3 | 3.5 | 3 | 4 | 3 | 0 | 0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | 0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 3 | 1 | 2
-25 | 0 | OutOfRng | 1 | OutOfRng | 0 | OutOfRng | OutOfRng | 0 | 0 | 0 | 0 | 2 | 2 | 0 | 2 | 3 | 3 | 3 | 3 | 4 | 3.5 | 4 | 3 | 3.5 | 3.5 | 3.5 | 4 | 3 | 3.5 | 3 | 4 | 3 | 0 | 0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | 0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 5 | 1 | 2
-19.99 | 0 | OutOfRng | OutOfRng | OutOfRng | 0 | OutOfRng | OutOfRng | 0 | 0 | 0 | 0 | OutOfRng | OutOfRng | 0 | OutOfRng | 3 | 3 | 3 | 3 | 4 | 3.5 | 4 | 3 | 3.5 | 3.5 | 3.5 | 4 | 3 | 3.5 | 3 | 4 | 3 | 0 | 0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | 0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OutOfRng | 0 | OutOfRng | 1 | 2
-15 | 0 | OutOfRng | OutOfRng | OutOfRng | 0 | OutOfRng | OutOfRng | 0 | 0 | 0 | 0 | OutOfRng | OutOfRng | 0 | OutOfRng | 3 | 3 | 3 | 3 | 4 | 3.5 | 4 | 3 | 3.5 | 3.5 | 3.5 | 4 | 3 | 3.5 | 3 | 4 | 3 | 0 | 0 | OutOfRng | OutOfRng | 0 | OutOfRng | 0 | 0 | OutOfRng | OutOfRng | 0 | OutOfRng | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OutOfRng | 0 | OutOfRng | 1 | 2
-10 | 0 | OutOfRng | OutOfRng | 0 | 0 | OutOfRng | 0 | 0 | 0 | 0 | 0 | OutOfRng | OutOfRng | 0 | OutOfRng | 4.5 | 4.5 | 2.5 | 2.5 | 4 | 3.5 | 4 | 2.5 | 3.5 | 3.5 | 3.5 | 4 | 2.5 | 3.5 | 2.5 | 4 | 2.5 | 0 | 0 | OutOfRng | OutOfRng | 0 | 0 | 0 | 0 | OutOfRng | OutOfRng | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OutOfRng | 0 | OutOfRng | 1 | 2
-5 | 0 | OutOfRng | OutOfRng | 0 | 0 | OutOfRng | 0 | 0 | 0 | 0 | 0 | OutOfRng | OutOfRng | 0 | OutOfRng | 4.5 | 4.5 | 2.5 | 2.5 | 4 | 3.5 | 4 | 2.5 | 3.5 | 3.5 | 3.5 | 4 | 2.5 | 3.5 | 2.5 | 4 | 2.5 | 0 | 0 | 0 | OutOfRng | 0 | 0 | 0 | 0 | 0 | OutOfRng | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OutOfRng | 0 | OutOfRng | 1 | 2
-1 | 0 | OutOfRng | OutOfRng | 0 | 0 | OutOfRng | 0 | 0 | 0 | 0 | 0 | OutOfRng | OutOfRng | 0 | OutOfRng | 4.5 | 4.5 | 2.5 | 2.5 | 4 | 3.5 | 4 | 2.5 | 3.5 | 3.5 | 3.5 | 4 | 2.5 | 3.5 | 2.5 | 4 | 2.5 | 0.52 | 0.52 | 0 | OutOfRng | 0 | 0 | 0.52 | 0.52 | 0 | OutOfRng | 0 | 0 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0 | OutOfRng | 0 | OutOfRng | 1 | 2
0 | 0 | 0 | OutOfRng | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OutOfRng | OutOfRng | 0 | OutOfRng | 4.5 | 4.5 | 2 | 2 | 3.5 | 3.5 | 3.5 | 2 | 3.5 | 3.5 | 3.5 | 3.5 | 2 | 3.5 | 2 | 3.5 | 2 | 0.52 | 0.52 | 0 | 0 | 0 | 0 | 0.52 | 0.52 | 0 | 0 | 0 | 0 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0 | OutOfRng | 0 | OutOfRng | 1 | 2
10 | 3 | 0 | OutOfRng | 0 | 3 | 0 | 0 | 3 | 3 | 3 | 3 | OutOfRng | OutOfRng | 3 | OutOfRng | 4.5 | 4.5 | 2 | 2 | 3 | 3.5 | 3 | 2 | 3.5 | 3.5 | 3.5 | 3 | 2 | 3.5 | 2 | 3 | 2 | 0.52 | 0.52 | 0 | 0 | 0 | 0 | 0.52 | 0.52 | 0 | 0 | 0 | 0 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | 3 | OutOfRng | 3 | OutOfRng | 2 | 2
15 | 3 | 0 | OutOfRng | 0 | 3 | 0 | 0 | 3 | 3 | 3 | 3 | OutOfRng | OutOfRng | 3 | OutOfRng | 4.5 | 4.5 | 2 | 2 | 3 | 3.5 | 3 | 2 | 3.5 | 3.5 | 3.5 | 3 | 2 | 3.5 | 2 | 3 | 2 | 0.8 | 0.8 | 0 | 0 | 0 | 0 | 0.8 | 0.8 | 0 | 0 | 0 | 0 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 3 | OutOfRng | 3 | OutOfRng | 2 | 2
17 | 3 | 0 | OutOfRng | 0 | 3 | 0 | 0 | 3 | 3 | 3 | 3 | OutOfRng | OutOfRng | 3 | OutOfRng | 4.5 | 4.5 | 2 | 2 | 3 | 3.5 | 3 | 2 | 3.5 | 3.5 | 3.5 | 3 | 2 | 3.5 | 2 | 3 | 2 | 0.8 | 0.8 | 0 | 0 | 0.77 | 0.77 | 0.8 | 0.8 | 0 | 0 | 0.77 | 0.77 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 3 | OutOfRng | 3 | OutOfRng | 2 | 2
20.01 | OutOfRng | 0 | OutOfRng | 0 | OutOfRng | 0 | 0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3 | 3 | 3 | OutOfRng | 3 | 3 | 3 | 3 | OutOfRng | 3 | OutOfRng | 3 | OutOfRng | OutOfRng | OutOfRng | 0 | 0 | 0.77 | 0.77 | OutOfRng | OutOfRng | 0 | 0 | 0.77 | 0.77 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
25 | OutOfRng | 0 | OutOfRng | 5 | OutOfRng | 0 | 5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3 | 3 | 3 | OutOfRng | 3 | 3 | 3 | 3 | OutOfRng | 3 | OutOfRng | 3 | OutOfRng | OutOfRng | OutOfRng | 0 | 0 | 0.77 | 0.77 | OutOfRng | OutOfRng | 0 | 0 | 0.77 | 0.77 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
29 | OutOfRng | 0 | OutOfRng | 5 | OutOfRng | 0 | 5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3 | 3 | 3 | OutOfRng | 3 | 3 | 3 | 3 | OutOfRng | 3 | OutOfRng | 3 | OutOfRng | OutOfRng | OutOfRng | 0.52 | 0.52 | 0.77 | 0.77 | OutOfRng | OutOfRng | 0.52 | 0.52 | 0.77 | 0.77 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
30 | OutOfRng | 0 | OutOfRng | 5 | OutOfRng | 0 | 5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3 | 3 | 3 | OutOfRng | 3 | 3 | 3 | 3 | OutOfRng | 3 | OutOfRng | 3 | OutOfRng | OutOfRng | OutOfRng | 0.52 | 0.52 | 2.84 | 2.84 | OutOfRng | OutOfRng | 0.52 | 0.52 | 2.84 | 2.84 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
30.01 | OutOfRng | 0 | OutOfRng | 5 | OutOfRng | 0 | 5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.5 | OutOfRng | OutOfRng | 2.5 | 2.5 | 2.5 | OutOfRng | OutOfRng | 2.5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0.52 | 0.52 | 2.84 | 2.84 | OutOfRng | OutOfRng | 0.52 | 0.52 | 2.84 | 2.84 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
35 | OutOfRng | 4 | OutOfRng | 5 | OutOfRng | 4 | 5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.5 | OutOfRng | OutOfRng | 2.5 | 2.5 | 2.5 | OutOfRng | OutOfRng | 2.5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0.52 | 0.52 | 2.84 | 2.84 | OutOfRng | OutOfRng | 0.52 | 0.52 | 2.84 | 2.84 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
35.01 | OutOfRng | 4 | OutOfRng | OutOfRng | OutOfRng | 4 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.5 | OutOfRng | OutOfRng | 2.5 | 2.5 | 2.5 | OutOfRng | OutOfRng | 2.5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0.52 | 0.52 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0.52 | 0.52 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
39 | OutOfRng | 4 | OutOfRng | OutOfRng | OutOfRng | 4 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.5 | OutOfRng | OutOfRng | 2.5 | 2.5 | 2.5 | OutOfRng | OutOfRng | 2.5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.66 | 1.66 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.66 | 1.66 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
44.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.5 | 2.5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
47.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
50.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
Input VSWR
Sensor
Freq (GHz) | 8481A | 8481B | 8481D | 8481H | 8482A | 8482B | 8482H | 8483A | 8483A-H03 | 8485A | 8485A-033 | 8485D | 8485D-033 | 8487A | 8487D | E4412A | E4413A | E9300A | E9300A-H24 | E9300A-H25 | E9300B | E9300H | E9301A | E9301B | E9301B-H01 | E9301B-H50 | E9301H | E9304A | E9304A-H01 | E9304A-H18 | E9304A-H19 | E9304A-H20 | N8481A | N8481A-CFT | N8481B | N8481B-CFT | N8481H | N8481H-CFT | N8482A | N8482A-CFT | N8482B | N8482B-CFT | N8482H | N8482H-CFT | N8485A | N8485A-033 | N8485A-033CFT | N8485A-CFT | N8486AQ | N8486AQ-CFT | N8486AR | N8486AR-CFT | N8487A | N8487A-CFT | N8488A | Q8486A | Q8486D | R8486A | R8486D | V8486A | W8486A
0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
9e-05 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.13 | 1.13 | 1.13 | 1.15 | 1.13 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.0001 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.6 | 1.1 | 1.2 | 1.8 | 1.8 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.13 | 1.13 | 1.13 | 1.15 | 1.13 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.54 | 1.54 | 1.08 | 1.08 | 1.13 | 1.13 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.0003 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.2 | 1.1 | 1.2 | 1.8 | 1.8 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.13 | 1.13 | 1.13 | 1.15 | 1.13 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.17 | 1.17 | 1.08 | 1.08 | 1.13 | 1.13 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.0006 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.2 | 1.1 | 1.2 | 1.18 | 1.18 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.13 | 1.13 | 1.13 | 1.15 | 1.13 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.17 | 1.17 | 1.08 | 1.08 | 1.13 | 1.13 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.001 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.1 | 1.1 | 1.2 | 1.18 | 1.18 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.13 | 1.13 | 1.13 | 1.15 | 1.13 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.06 | 1.06 | 1.08 | 1.08 | 1.13 | 1.13 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.01 | 1.4 | 1.1 | 1.4 | 1.2 | 1.1 | 1.1 | 1.2 | 1.18 | 1.18 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.12 | OutOfRng | 1.15 | OutOfRng | OutOfRng | 1.12 | 1.15 | 1.15 | 1.12 | 1.2 | 1.12 | 1.15 | 1.13 | 1.13 | 1.13 | 1.15 | 1.13 | 1.37 | 1.37 | 1.09 | 1.09 | 1.16 | 1.16 | 1.06 | 1.06 | 1.08 | 1.08 | 1.13 | 1.13 | 1.33 | 1.33 | 1.33 | 1.33 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.03 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.03 | 1.18 | 1.1 | 1.15 | 1.2 | 1.1 | 1.1 | 1.2 | 1.18 | 1.18 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.12 | OutOfRng | 1.13 | OutOfRng | OutOfRng | 1.12 | 1.15 | 1.13 | 1.12 | 1.2 | 1.12 | 1.15 | 1.13 | 1.13 | 1.13 | 1.15 | 1.13 | 1.14 | 1.14 | 1.09 | 1.09 | 1.16 | 1.16 | 1.06 | 1.06 | 1.08 | 1.08 | 1.13 | 1.13 | 1.33 | 1.33 | 1.33 | 1.33 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.03 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.04 | 1.18 | 1.1 | 1.15 | 1.2 | 1.1 | 1.1 | 1.2 | 1.18 | 1.18 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.15 | OutOfRng | 1.13 | OutOfRng | OutOfRng | 1.12 | 1.15 | 1.13 | 1.12 | 1.2 | 1.12 | 1.15 | 1.13 | 1.13 | 1.13 | 1.15 | 1.13 | 1.14 | 1.14 | 1.09 | 1.09 | 1.16 | 1.16 | 1.06 | 1.06 | 1.08 | 1.08 | 1.13 | 1.13 | 1.33 | 1.33 | 1.33 | 1.33 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.03 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.05 | 1.1 | 1.1 | 1.15 | 1.2 | 1.1 | 1.1 | 1.2 | 1.18 | 1.18 | 1.15 | 1.15 | 1.19 | 1.19 | 1.15 | 1.19 | 1.15 | 1.21 | 1.13 | 1.13 | 1.13 | 1.12 | 1.15 | 1.13 | 1.12 | 1.2 | 1.12 | 1.15 | 1.13 | 1.13 | 1.13 | 1.15 | 1.13 | 1.08 | 1.08 | 1.09 | 1.09 | 1.16 | 1.16 | 1.06 | 1.06 | 1.08 | 1.08 | 1.13 | 1.13 | 1.08 | 1.08 | 1.08 | 1.08 | 1.4 | 1.4 | 1.4 | 1.4 | 1.08 | 1.08 | 1.03 | 1.4 | 1.4 | 1.4 | 1.4 | 1.4 | 1.4
0.1 | 1.1 | 1.1 | 1.15 | 1.2 | 1.1 | 1.1 | 1.2 | 1.18 | 1.18 | 1.1 | 1.1 | 1.15 | 1.15 | 1.1 | 1.15 | 1.15 | 1.19 | 1.13 | 1.13 | 1.13 | 1.12 | 1.15 | 1.13 | 1.12 | 1.2 | 1.12 | 1.15 | 1.13 | 1.13 | 1.13 | 1.15 | 1.13 | 1.08 | 1.08 | 1.09 | 1.09 | 1.16 | 1.16 | 1.06 | 1.06 | 1.08 | 1.08 | 1.13 | 1.13 | 1.05 | 1.05 | 1.05 | 1.05 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.05 | 1.05 | 1.04 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
2.01 | 1.18 | 1.18 | 1.15 | 1.2 | 1.3 | 1.18 | 1.2 | OutOfRng | 1.18 | 1.15 | 1.15 | 1.15 | 1.15 | 1.15 | 1.15 | 1.17 | 1.19 | 1.19 | 1.19 | 1.19 | 1.12 | 1.15 | 1.19 | 1.12 | 1.3 | 1.12 | 1.15 | 1.19 | 1.19 | 1.19 | 1.15 | 1.19 | 1.16 | 1.16 | 1.14 | 1.14 | 1.16 | 1.16 | 1.07 | 1.07 | 1.16 | 1.16 | 1.13 | 1.13 | 1.14 | 1.14 | 1.14 | 1.14 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.1 | 1.1 | 1.04 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
2.41 | 1.18 | 1.18 | 1.15 | 1.2 | 1.3 | 1.18 | 1.2 | OutOfRng | 1.18 | 1.15 | 1.15 | 1.15 | 1.15 | 1.15 | 1.15 | 1.17 | 1.19 | 1.19 | 1.19 | 1.19 | 1.12 | 1.15 | 1.19 | 1.12 | 1.3 | 1.12 | 1.15 | 1.19 | 1.19 | 1.19 | 1.15 | 1.19 | 1.16 | 1.16 | 1.14 | 1.14 | 1.16 | 1.16 | 1.07 | 1.07 | 1.16 | 1.16 | 1.13 | 1.13 | 1.14 | 1.14 | 1.14 | 1.14 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.1 | 1.1 | 1.09 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
3.01 | 1.18 | 1.18 | 1.15 | 1.2 | 1.3 | 1.18 | 1.2 | OutOfRng | OutOfRng | 1.15 | 1.15 | 1.15 | 1.15 | 1.15 | 1.15 | 1.17 | 1.19 | 1.19 | 1.19 | 1.19 | 1.12 | 1.15 | 1.19 | 1.12 | 1.4 | 1.12 | 1.15 | 1.19 | 1.19 | 1.19 | 1.15 | 1.19 | 1.16 | 1.16 | 1.14 | 1.14 | 1.16 | 1.16 | 1.07 | 1.07 | 1.16 | 1.16 | 1.13 | 1.13 | 1.14 | 1.14 | 1.14 | 1.14 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.1 | 1.1 | 1.09 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
4.01 | 1.18 | 1.18 | 1.2 | 1.2 | 1.3 | 1.18 | 1.2 | OutOfRng | OutOfRng | 1.15 | 1.15 | 1.19 | 1.19 | 1.15 | 1.2 | 1.17 | 1.19 | 1.19 | 1.19 | 1.19 | 1.12 | 1.15 | 1.19 | 1.12 | 1.4 | 1.12 | 1.15 | 1.19 | 1.19 | 1.19 | 1.15 | 1.19 | 1.16 | 1.16 | 1.14 | 1.14 | 1.16 | 1.16 | 1.07 | 1.07 | 1.16 | 1.16 | 1.13 | 1.13 | 1.14 | 1.14 | 1.14 | 1.14 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.1 | 1.1 | 1.09 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
4.21 | 1.18 | 1.18 | 1.2 | 1.2 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.15 | 1.15 | 1.19 | 1.19 | 1.15 | 1.2 | 1.17 | 1.19 | 1.19 | 1.19 | 1.19 | 1.12 | 1.15 | 1.19 | 1.12 | OutOfRng | 1.12 | 1.15 | 1.19 | 1.19 | 1.19 | 1.15 | 1.19 | 1.16 | 1.16 | 1.14 | 1.14 | 1.16 | 1.16 | 1.07 | 1.07 | 1.16 | 1.16 | 1.13 | 1.13 | 1.14 | 1.14 | 1.14 | 1.14 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.1 | 1.1 | 1.09 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
6.01 | 1.18 | 1.18 | 1.2 | 1.2 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.15 | 1.15 | 1.19 | 1.19 | 1.15 | 1.2 | 1.2 | 1.19 | 1.19 | 1.19 | 1.19 | 1.12 | 1.15 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.19 | 1.19 | 1.15 | 1.19 | 1.16 | 1.16 | 1.14 | 1.14 | 1.16 | 1.16 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.14 | 1.14 | 1.14 | 1.14 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.1 | 1.1 | 1.09 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
8.01 | 1.18 | 1.18 | 1.2 | 1.25 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.15 | 1.15 | 1.19 | 1.19 | 1.15 | 1.2 | 1.2 | 1.21 | 1.19 | 1.19 | 1.19 | 1.17 | 1.25 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.19 | 1.19 | 1.25 | 1.19 | 1.16 | 1.16 | 1.14 | 1.14 | 1.22 | 1.22 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.14 | 1.14 | 1.14 | 1.14 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.1 | 1.1 | 1.09 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
10.01 | 1.18 | 1.18 | 1.3 | 1.25 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.15 | 1.15 | 1.19 | 1.19 | 1.15 | 1.2 | 1.2 | 1.21 | 1.19 | 1.19 | 1.19 | 1.17 | 1.25 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.19 | 1.19 | 1.25 | 1.19 | 1.16 | 1.16 | 1.14 | 1.14 | 1.22 | 1.22 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.14 | 1.14 | 1.14 | 1.14 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.1 | 1.1 | 1.09 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
11.01 | 1.18 | 1.18 | 1.3 | 1.25 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.15 | 1.15 | 1.19 | 1.19 | 1.15 | 1.2 | 1.27 | 1.21 | 1.19 | 1.19 | 1.19 | 1.17 | 1.25 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.19 | 1.19 | 1.25 | 1.19 | 1.16 | 1.16 | 1.14 | 1.14 | 1.22 | 1.22 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.14 | 1.14 | 1.14 | 1.14 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.1 | 1.1 | 1.09 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
12.01 | 1.18 | 1.18 | 1.3 | 1.25 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.15 | 1.15 | 1.25 | 1.25 | 1.15 | 1.2 | 1.27 | 1.21 | 1.19 | 1.19 | 1.19 | 1.17 | 1.25 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.19 | 1.19 | 1.25 | 1.19 | 1.16 | 1.16 | 1.14 | 1.14 | 1.22 | 1.22 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.14 | 1.14 | 1.14 | 1.14 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.1 | 1.1 | 1.09 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
12.41 | 1.28 | 1.28 | 1.3 | 1.3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.2 | 1.2 | 1.25 | 1.25 | 1.2 | 1.29 | 1.27 | 1.21 | 1.19 | 1.19 | 1.19 | 1.24 | 1.28 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.22 | 1.22 | 1.28 | 1.22 | 1.23 | 1.23 | 1.23 | 1.23 | 1.32 | 1.32 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.19 | 1.19 | 1.19 | 1.19 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.16 | 1.16 | 1.12 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
14.01 | 1.28 | 1.28 | 1.3 | 1.3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.2 | 1.2 | 1.25 | 1.25 | 1.2 | 1.29 | 1.27 | 1.21 | 1.22 | 1.22 | 1.25 | 1.24 | 1.28 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.22 | 1.22 | 1.28 | 1.22 | 1.23 | 1.23 | 1.23 | 1.23 | 1.32 | 1.32 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.19 | 1.19 | 1.19 | 1.19 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.16 | 1.16 | 1.12 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
15.01 | 1.28 | 1.28 | 1.5 | 1.3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.2 | 1.2 | 1.25 | 1.25 | 1.2 | 1.29 | 1.27 | 1.21 | 1.22 | 1.22 | 1.25 | 1.24 | 1.28 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.22 | 1.22 | 1.28 | 1.22 | 1.23 | 1.23 | 1.23 | 1.23 | 1.32 | 1.32 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.19 | 1.19 | 1.19 | 1.19 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.16 | 1.16 | 1.12 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
16.01 | 1.28 | 1.28 | 1.5 | 1.3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.2 | 1.2 | 1.25 | 1.25 | 1.2 | 1.29 | 1.27 | 1.21 | 1.26 | 1.26 | 1.26 | 1.24 | 1.28 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.27 | 1.27 | 1.28 | 1.27 | 1.23 | 1.23 | 1.23 | 1.23 | 1.32 | 1.32 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.19 | 1.19 | 1.19 | 1.19 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.16 | 1.16 | 1.12 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
18.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.25 | 1.25 | 1.29 | 1.29 | 1.25 | 1.37 | OutOfRng | 1.26 | OutOfRng | 1.3 | 1.3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.26 | 1.26 | 1.26 | 1.26 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.22 | 1.22 | 1.22 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
24 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.25 | 1.25 | 1.29 | 1.29 | 1.25 | 1.37 | OutOfRng | 1.26 | OutOfRng | 1.3 | 1.3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.26 | 1.26 | 1.26 | 1.26 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.22 | 1.22 | 1.22 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
24 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.25 | 1.25 | 1.29 | 1.29 | 1.25 | 1.37 | OutOfRng | 1.26 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.26 | 1.26 | 1.26 | 1.26 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.22 | 1.22 | 1.22 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
26.5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.25 | 1.25 | 1.29 | 1.29 | 1.25 | 1.37 | OutOfRng | 1.26 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.26 | 1.26 | 1.26 | 1.26 | OutOfRng | OutOfRng | 1.4 | 1.4 | 1.22 | 1.22 | 1.22 | OutOfRng | OutOfRng | 1.4 | 1.4 | OutOfRng | OutOfRng
26.51 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.4 | OutOfRng | 1.35 | 1.3 | 1.37 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.32 | 1.32 | OutOfRng | OutOfRng | OutOfRng | 1.4 | 1.4 | 1.3 | 1.3 | 1.31 | OutOfRng | OutOfRng | 1.4 | 1.4 | OutOfRng | OutOfRng
33 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.4 | OutOfRng | 1.35 | 1.3 | 1.37 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.32 | 1.32 | OutOfRng | 1.5 | 1.5 | 1.4 | 1.4 | 1.3 | 1.3 | 1.31 | 1.5 | 1.4 | 1.4 | 1.4 | OutOfRng | OutOfRng
33.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.3 | 1.37 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.5 | 1.5 | 1.4 | 1.4 | 1.3 | 1.3 | 1.31 | 1.5 | 1.4 | 1.4 | 1.4 | OutOfRng | OutOfRng
34.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.3 | 1.61 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.5 | 1.5 | 1.4 | 1.4 | 1.3 | 1.3 | 1.31 | 1.5 | 1.4 | 1.4 | 1.4 | OutOfRng | OutOfRng
40.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.5 | 1.89 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.5 | 1.5 | OutOfRng | OutOfRng | 1.34 | 1.34 | 1.47 | 1.5 | 1.4 | OutOfRng | OutOfRng | OutOfRng | OutOfRng
50 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.5 | 1.89 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.5 | 1.5 | OutOfRng | OutOfRng | 1.34 | 1.34 | 1.47 | 1.5 | 1.4 | OutOfRng | OutOfRng | 1.06 | OutOfRng
50.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.47 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.06 | OutOfRng
67.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.06 | OutOfRng
70.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.06 | OutOfRng
75 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.06 | 1.08
75.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.08
110.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
Calibration Factor (%)
Sensor
Freq (GHz) | 8481A | 8481B | 8481D | 8481H | 8482A | 8482B | 8482H | 8483A | 8483A-H03 | 8485A | 8485A-033 | 8485D | 8485D-033 | 8487A | 8487D | E4412A | E4413A | E9300AH24Hi | E9300AH24Lo | E9300AH25Hi | E9300AH25Lo | E9300AHi | E9300ALo | E9300BHi | E9300BLo | E9300HHi | E9300HLo | E9301AHi | E9301ALo | E9301BH01Hi | E9301BH01Lo | E9301BH50Hi | E9301BH50Lo | E9301BHi | E9301BLo | E9301HHi | E9301HLo | E9304AH01Hi | E9304AH01Lo | E9304AH18Hi | E9304AH18Lo | E9304AH19Hi | E9304AH19Lo | E9304AH20Hi | E9304AH20Lo | E9304AHi | E9304ALo | N8481ACFT | N8481AHi | N8481ALo | N8481BCFT | N8481BHi | N8481BLo | N8481HCFT | N8481HHi | N8481HLo | N8482ACFT | N8482AHi | N8482ALo | N8482BCFT | N8482BHi | N8482BLo | N8482HCFT | N8482HHi | N8482HLo | N8485A033CFT | N8485A033Hi | N8485A033Lo | N8485ACFT | N8485AHi | N8485ALo | N8486AQCFT | N8486AQHi | N8486AQLo | N8486ARCFT | N8486ARHi | N8486ARLo | N8487ACFT | N8487AHi | N8487ALo | N8488AHi | N8488ALo | Q8486A | Q8486D | R8486A | R8486D | V8486A | W8486A
0 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
9e-05 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2 | 1.7 | 2 | 1.7 | 2.3 | 1.7 | 2 | 1.7 | 1.8 | 1.6 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.0001 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0.87 | 1.5 | 0.91 | 1.59 | 1.59 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2 | 1.7 | 2 | 1.7 | 2.3 | 1.7 | 2 | 1.7 | 1.8 | 1.6 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 0.91 | 0.91 | 0.91 | 1.48 | 1.48 | 1.48 | 0.89 | 0.89 | 0.89 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.01 | 0.81 | 1.48 | 0.77 | 0.86 | 0.8 | 1.44 | 0.81 | 1.39 | 1.39 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.8 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.1 | 1.8 | 2.1 | 1.8 | 2.1 | 1.8 | 2.1 | 1.8 | 2.1 | 1.8 | 2.1 | 1.8 | 2.1 | 1.8 | 2.1 | 1.8 | 2 | 1.7 | 2 | 1.7 | 2.3 | 1.7 | 2 | 1.7 | 1.8 | 1.6 | 0.82 | 0.82 | 0.82 | 1.42 | 1.42 | 1.42 | 0.77 | 0.77 | 0.77 | 0.78 | 0.78 | 0.78 | 1.43 | 1.43 | 1.43 | 0.79 | 0.79 | 0.79 | 0.82 | 0.82 | 0.82 | 0.82 | 0.82 | 0.82 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.7 | 1.7 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.03 | 0.78 | 1.48 | 0.77 | 0.86 | 0.8 | 1.44 | 0.81 | 1.39 | 1.39 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.8 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.1 | 1.8 | 2.1 | 1.8 | 2.1 | 1.8 | 2.1 | 1.8 | 2.1 | 1.8 | 2.1 | 1.8 | 2.1 | 1.8 | 2.1 | 1.8 | 2 | 1.7 | 2 | 1.7 | 2.3 | 1.7 | 2 | 1.7 | 1.8 | 1.6 | 0.77 | 0.77 | 0.82 | 1.42 | 1.42 | 1.42 | 0.77 | 0.77 | 0.77 | 0.77 | 0.77 | 0.77 | 1.43 | 1.43 | 1.43 | 0.79 | 0.79 | 0.79 | 0.82 | 0.82 | 0.82 | 0.82 | 0.82 | 0.82 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.7 | 1.7 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.031 | 0.78 | 1.48 | 0.81 | 0.88 | 0.8 | 1.51 | 0.92 | 1.41 | 1.41 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.8 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 2 | 1.7 | 2 | 1.7 | 2.3 | 1.7 | 2 | 1.7 | 1.8 | 1.6 | 0.77 | 0.77 | 0.77 | 1.48 | 1.48 | 1.48 | 0.89 | 0.89 | 0.89 | 0.77 | 0.77 | 0.77 | 1.49 | 1.49 | 1.49 | 0.89 | 0.89 | 0.89 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.7 | 1.7 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.05 | 0.78 | 1.48 | 0.81 | 0.88 | 0.8 | 1.51 | 0.92 | 1.41 | 1.41 | 1.29 | 1.29 | 1.24 | 1.24 | 1.38 | 1.33 | 1.8 | OutOfRng | 2.3 | 1.6 | 2.3 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 2 | 1.7 | 2 | 1.7 | 2.3 | 1.7 | 2 | 1.7 | 1.8 | 1.6 | 0.77 | 0.77 | 0.77 | 1.48 | 1.48 | 1.48 | 0.89 | 0.89 | 0.89 | 0.77 | 0.77 | 0.77 | 1.49 | 1.49 | 1.49 | 0.89 | 0.89 | 0.89 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.33 | 1.33 | 1.33 | 1.6 | 1.6 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.1 | 0.78 | 1.48 | 0.81 | 0.88 | 0.8 | 1.51 | 0.92 | 1.41 | 1.41 | 1.29 | 1.29 | 1.24 | 1.24 | 1.38 | 1.33 | 1.8 | OutOfRng | 2.3 | 1.6 | 2.3 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 2 | 1.7 | 2 | 1.7 | 2.3 | 1.7 | 2 | 1.7 | 1.8 | 1.6 | 0.77 | 0.77 | 0.77 | 1.48 | 1.48 | 1.48 | 0.89 | 0.89 | 0.89 | 0.77 | 0.77 | 0.77 | 1.49 | 1.49 | 1.49 | 0.89 | 0.89 | 0.89 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.33 | 1.33 | 1.33 | 1.58 | 1.58 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.5 | 0.78 | 1.48 | 0.81 | 0.87 | 0.8 | 1.48 | 0.88 | 1.41 | 1.41 | 1.26 | 1.26 | 1.24 | 1.24 | 1.34 | 1.33 | 1.8 | 1.8 | 2.3 | 1.6 | 2.3 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 1.8 | 1.6 | 2 | 1.7 | 2 | 1.7 | 2.3 | 1.7 | 2 | 1.7 | 1.8 | 1.6 | 0.77 | 0.77 | 0.77 | 1.48 | 1.48 | 1.48 | 0.89 | 0.89 | 0.89 | 0.77 | 0.77 | 0.77 | 1.49 | 1.49 | 1.49 | 0.89 | 0.89 | 0.89 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.33 | 1.33 | 1.33 | 1.58 | 1.58 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
0.51 | 0.78 | 1.48 | 0.81 | 0.87 | 0.8 | 1.48 | 0.88 | 1.41 | 1.41 | 1.26 | 1.26 | 1.26 | 1.26 | 1.34 | 1.35 | 1.8 | 1.8 | 2.8 | 1.8 | 2.8 | 1.8 | 2.3 | 1.8 | 2.3 | 1.8 | 2.3 | 1.8 | 2.3 | 1.8 | 2.3 | 1.8 | 2.3 | 1.8 | 2.3 | 1.8 | 2.3 | 1.8 | 2.2 | 1.7 | 2.2 | 1.7 | 2.8 | 1.7 | 2.2 | 1.7 | 2.3 | 1.8 | 0.78 | 0.78 | 0.78 | 1.48 | 1.48 | 1.48 | 0.89 | 0.89 | 0.89 | 0.78 | 0.78 | 0.78 | 1.49 | 1.49 | 1.49 | 0.89 | 0.89 | 0.89 | 1.26 | 1.26 | 1.26 | 1.26 | 1.26 | 1.26 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.35 | 1.35 | 1.35 | 1.58 | 1.58 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
1.2 | 0.78 | 1.48 | 0.81 | 0.87 | 0.8 | 1.48 | 0.88 | 1.41 | 1.41 | 1.26 | 1.26 | 1.26 | 1.26 | 1.34 | 1.35 | 1.8 | 1.8 | 2.3 | 1.7 | 2.3 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 2.2 | 1.7 | 2.2 | 1.7 | 2.8 | 1.7 | 2.2 | 1.7 | 1.8 | 1.7 | 0.78 | 0.78 | 0.78 | 1.48 | 1.48 | 1.48 | 0.89 | 0.89 | 0.89 | 0.78 | 0.78 | 0.78 | 1.49 | 1.49 | 1.49 | 0.89 | 0.89 | 0.89 | 1.26 | 1.26 | 1.26 | 1.26 | 1.26 | 1.26 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.35 | 1.35 | 1.35 | 1.58 | 1.58 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
1.21 | 0.92 | 1.54 | 0.97 | 0.98 | 0.9 | 1.53 | 0.95 | 1.46 | 1.46 | 1.35 | 1.35 | 1.35 | 1.35 | 1.41 | 1.41 | 1.8 | 1.8 | 2.3 | 1.7 | 2.3 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 2.2 | 1.7 | 2.2 | 1.7 | 2.3 | 1.7 | 2.2 | 1.7 | 1.8 | 1.7 | 0.91 | 0.91 | 0.91 | 1.58 | 1.58 | 1.58 | 1.06 | 1.06 | 1.06 | 0.89 | 0.89 | 0.89 | 1.56 | 1.56 | 1.56 | 1.02 | 1.02 | 1.02 | 1.35 | 1.35 | 1.35 | 1.35 | 1.35 | 1.35 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.41 | 1.41 | 1.41 | 1.58 | 1.58 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
2.01 | 0.92 | 1.54 | 0.97 | 0.98 | 0.9 | 1.53 | 0.95 | OutOfRng | 2 | 1.35 | 1.35 | 1.35 | 1.35 | 1.41 | 1.41 | 2.4 | 2.4 | 2.3 | 1.7 | 2.3 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 2.3 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 0.91 | 0.91 | 0.91 | 1.58 | 1.58 | 1.58 | 1.06 | 1.06 | 1.06 | 0.89 | 0.89 | 0.89 | 1.56 | 1.56 | 1.56 | 1.02 | 1.02 | 1.02 | 1.35 | 1.35 | 1.35 | 1.35 | 1.35 | 1.35 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.41 | 1.41 | 1.41 | 1.75 | 1.75 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
3.01 | 0.92 | 1.54 | 0.97 | 0.98 | 0.9 | 1.53 | 0.95 | OutOfRng | OutOfRng | 1.35 | 1.35 | 1.35 | 1.35 | 1.41 | 1.41 | 2.4 | 2.4 | 2.3 | 1.7 | 2.3 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 2.3 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 0.91 | 0.91 | 0.91 | 1.58 | 1.58 | 1.58 | 1.06 | 1.06 | 1.06 | 0.89 | 0.89 | 0.89 | 1.56 | 1.56 | 1.56 | 1.02 | 1.02 | 1.02 | 1.35 | 1.35 | 1.35 | 1.35 | 1.35 | 1.35 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.41 | 1.41 | 1.41 | 1.75 | 1.75 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
4.21 | 0.92 | 1.54 | 0.97 | 0.98 | 0.9 | 1.53 | 0.95 | OutOfRng | OutOfRng | 1.35 | 1.35 | 1.35 | 1.35 | 1.41 | 1.41 | 2.4 | 2.4 | 2.3 | 1.7 | 2.3 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | OutOfRng | OutOfRng | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 2.3 | 1.7 | 1.8 | 1.7 | 1.8 | 1.7 | 0.91 | 0.91 | 0.91 | 1.58 | 1.58 | 1.58 | 1.06 | 1.06 | 1.06 | 0.89 | 0.89 | 0.89 | 1.56 | 1.56 | 1.56 | 1.02 | 1.02 | 1.02 | 1.35 | 1.35 | 1.35 | 1.35 | 1.35 | 1.35 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.41 | 1.41 | 1.41 | 1.75 | 1.75 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
6.01 | 1.16 | 1.71 | 1.2 | 1.36 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.61 | 1.61 | 1.63 | 1.63 | 1.59 | 1.62 | 2.4 | 2.4 | 2.4 | 1.8 | 2.4 | 1.8 | 1.9 | 1.8 | 1.9 | 1.8 | 1.9 | 1.8 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.9 | 1.8 | 1.9 | 1.8 | 2.4 | 1.8 | 1.9 | 1.8 | OutOfRng | OutOfRng | 1.26 | 1.26 | 1.26 | 1.77 | 1.77 | 1.77 | 1.46 | 1.46 | 1.46 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.66 | 1.66 | 1.66 | 1.66 | 1.66 | 1.66 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.61 | 1.61 | 1.61 | 1.75 | 1.75 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
10.01 | 1.16 | 1.71 | 1.2 | 1.36 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.61 | 1.61 | 1.63 | 1.63 | 1.59 | 1.62 | 2.4 | 2.6 | 2.4 | 1.8 | 2.4 | 1.8 | 1.9 | 1.8 | 1.9 | 1.8 | 1.9 | 1.8 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.9 | 1.8 | 1.9 | 1.8 | 2.4 | 1.8 | 1.9 | 1.8 | OutOfRng | OutOfRng | 1.26 | 1.26 | 1.26 | 1.77 | 1.77 | 1.77 | 1.46 | 1.46 | 1.46 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.66 | 1.66 | 1.66 | 1.66 | 1.66 | 1.66 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.61 | 1.61 | 1.61 | 1.75 | 1.75 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
12.01 | 1.16 | 1.71 | 1.2 | 1.36 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.61 | 1.61 | 1.63 | 1.63 | 1.59 | 1.62 | 2.4 | 2.8 | 2.4 | 1.8 | 2.4 | 1.8 | 1.9 | 1.8 | 1.9 | 1.8 | 1.9 | 1.8 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.9 | 1.8 | 1.9 | 1.8 | 2.4 | 1.8 | 1.9 | 1.8 | OutOfRng | OutOfRng | 1.26 | 1.26 | 1.26 | 1.77 | 1.77 | 1.77 | 1.46 | 1.46 | 1.46 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.66 | 1.66 | 1.66 | 1.66 | 1.66 | 1.66 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.61 | 1.61 | 1.61 | 1.75 | 1.75 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
12.41 | 1.16 | 1.71 | 1.2 | 1.36 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.61 | 1.61 | 1.63 | 1.63 | 1.59 | 1.62 | 2.4 | 2.8 | 2.4 | 1.8 | 2.4 | 1.8 | 1.9 | 1.8 | 1.9 | 1.8 | 1.9 | 1.8 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.9 | 1.8 | 1.9 | 1.8 | 2.4 | 1.8 | 1.9 | 1.8 | OutOfRng | OutOfRng | 1.26 | 1.26 | 1.26 | 1.77 | 1.77 | 1.77 | 1.46 | 1.46 | 1.46 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.66 | 1.66 | 1.66 | 1.66 | 1.66 | 1.66 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.61 | 1.61 | 1.61 | 1.92 | 1.92 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
14.01 | 1.59 | 1.99 | 1.72 | 1.71 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.77 | 1.77 | 1.83 | 1.83 | 1.69 | 1.73 | 2.4 | 2.8 | 2.7 | 2 | 2.7 | 2 | 2.2 | 2 | 2.2 | 2 | 2.2 | 2 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.2 | 2 | 2.2 | 2 | 2.7 | 2 | 2.2 | 2 | OutOfRng | OutOfRng | 1.59 | 1.59 | 1.59 | 1.92 | 1.92 | 1.92 | 1.73 | 1.73 | 1.73 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.83 | 1.83 | 1.83 | 1.83 | 1.83 | 1.83 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.73 | 1.73 | 1.73 | 1.92 | 1.92 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
16.01 | 1.59 | 1.99 | 1.72 | 1.71 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.77 | 1.77 | 1.83 | 1.83 | 1.69 | 1.73 | 2.6 | 2.8 | 2.7 | 2 | 2.7 | 2 | 2.2 | 2 | 2.2 | 2 | 2.2 | 2 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.2 | 2 | 2.2 | 2 | 2.7 | 2 | 2.2 | 2 | OutOfRng | OutOfRng | 1.59 | 1.59 | 1.59 | 1.92 | 1.92 | 1.92 | 1.73 | 1.73 | 1.73 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.83 | 1.83 | 1.83 | 1.83 | 1.83 | 1.83 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 1.73 | 1.73 | 1.73 | 1.92 | 1.92 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
18.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.47 | 2.47 | 2.45 | 2.45 | 2.23 | 2.25 | OutOfRng | 2.8 | 3.5 | 3 | 3.5 | 3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.67 | 2.67 | 2.67 | 2.67 | 2.67 | 2.67 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.26 | 2.26 | 2.26 | 2.43 | 2.43 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
20.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.47 | 2.47 | 2.45 | 2.45 | 2.23 | 2.25 | OutOfRng | 3 | 3.5 | 3 | 3.5 | 3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.67 | 2.67 | 2.67 | 2.67 | 2.67 | 2.67 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.26 | 2.26 | 2.26 | 2.43 | 2.43 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
24.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.47 | 2.47 | 2.45 | 2.45 | 2.23 | 2.25 | OutOfRng | 3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.67 | 2.67 | 2.67 | 2.67 | 2.67 | 2.67 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.26 | 2.26 | 2.26 | 2.43 | 2.43 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
26.5 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.47 | 2.47 | 2.45 | 2.45 | 2.23 | 2.25 | OutOfRng | 3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.67 | 2.67 | 2.67 | 2.67 | 2.67 | 2.67 | OutOfRng | OutOfRng | OutOfRng | 2.68 | 2.68 | 2.68 | 2.26 | 2.26 | 2.26 | 2.43 | 2.43 | OutOfRng | OutOfRng | 2.79 | 2.47 | OutOfRng | OutOfRng
26.51 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3.35 | OutOfRng | 2.94 | 2.58 | 2.55 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3.32 | 3.32 | 3.32 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.68 | 2.68 | 2.68 | 2.58 | 2.58 | 2.58 | 4.56 | 4.56 | OutOfRng | OutOfRng | 2.79 | 2.47 | OutOfRng | OutOfRng
33 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3.35 | OutOfRng | 2.94 | 2.58 | 2.55 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3.32 | 3.32 | 3.32 | OutOfRng | OutOfRng | OutOfRng | 3.14 | 3.14 | 3.14 | 3.19 | 3.19 | 3.19 | 2.58 | 2.58 | 2.58 | 4.56 | 4.56 | 3.12 | 2.85 | 2.79 | 2.42 | OutOfRng | OutOfRng
33.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.73 | 3.08 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3.14 | 3.14 | 3.14 | 3.19 | 3.19 | 3.19 | 2.8 | 2.8 | 2.8 | 4.56 | 4.56 | 3.12 | 2.85 | 2.93 | 2.42 | OutOfRng | OutOfRng
34.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.73 | 3.08 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3.4 | 3.4 | 3.4 | 3.19 | 3.19 | 3.19 | 2.8 | 2.8 | 2.8 | 4.56 | 4.56 | 3.4 | 2.85 | 2.93 | 2.42 | OutOfRng | OutOfRng
35 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.73 | 3.08 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3.14 | 3.14 | 3.14 | 3.19 | 3.19 | 3.19 | 2.8 | 2.8 | 2.8 | 4.56 | 4.56 | 3.12 | 2.85 | 2.93 | 2.42 | OutOfRng | OutOfRng
40 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 2.73 | 3.08 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3.14 | 3.14 | 3.14 | 3.19 | 3.19 | 3.19 | 2.8 | 2.8 | 2.8 | 4.56 | 4.56 | 3.12 | 2.84 | 2.93 | 2.42 | OutOfRng | OutOfRng
40.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3.67 | 4.28 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3.19 | 3.19 | 3.19 | OutOfRng | OutOfRng | OutOfRng | 3.66 | 3.66 | 3.66 | 4.56 | 4.56 | 3.15 | 2.84 | OutOfRng | OutOfRng | OutOfRng | OutOfRng
45.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 4.33 | 4.72 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 3.26 | 3.26 | 3.26 | OutOfRng | OutOfRng | OutOfRng | 4.23 | 4.23 | 4.23 | 4.56 | 4.56 | 3.34 | 2.84 | OutOfRng | OutOfRng | OutOfRng | OutOfRng
50.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 5.3 | 5.3 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 6.7 | OutOfRng
67.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 5.81 | 5.81 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 6.7 | OutOfRng
68 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 5.81 | 5.81 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 7.3 | OutOfRng
70.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 7.3 | OutOfRng
75 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 5.1 | 5.5
75.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | 5.4
110.01 | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng | OutOfRng
Note on B5: Francis Small: Use the 25 +/- 10 C value.
Note on AE52: Francis Small: Not sure what the specs are - copied them from the E9300B
Note on AE94: Francis Small: Not sure what the spec is, so copied them from E9304A-H18
Note on AN124: Francis Small: Not sure what the specs are - it seemed logical to make them the same as E9304A H18.
Note on AN138: Francis Small: Grey range is copied from E9300B