COMPATIBLE  MODE & Address display, etc.

(6030A, 6031A, 6032A, 6033A, 6035A, 6038A)

(for SCPI, instruments must have the 4-pin INH-FLT plug on the rear panel)

These products are shipped from the factory in the Compatible Language Mode with their hardware address set to "Adr 5".

The programming language can only be changed to SCPI (TMSL) Mode via a controller by sending the string:

"SYST:LANG<space>TMSL"

In the SCPI Mode, either by cycling the ac line or holding in the [LCL] front panel button, the display will show  "Adr 5", then "SEC  xxx"

(SEC is the secondary address, available only in the SCPI mode)

With the [LCL] button depressed, and the display showing "SEC  xxx", the rotary pulse generator will change the xxx portion  from "---" (null) to a number between "0" and "15".

The display "Adr  5" followed by "SEC  ---" designates the unit to be a stand-alone unit and must be talked to by using address 705.

A display of "SEC 0" designates the unit as the direct unit in a possible chain of instruments and must be talked to by using address 70500.  (This unit must be the unit connected to the controller via the GP-IB cable.)

In either case ("SEC ---", or "SEC  0"), the display "Adr  5" is still viewable when the ac power is cycled or when the [LCL] button is initially depressed.

If the product has been set to "SEC  1", "SEC  2" or "SEC  nn"; when the ac line is cycled or the [LCL] button is initially depressed, the display will first show "Adr  ---" (null), then "SEC  nn" (the secondary address given to this specific unit).

An instrument with a "SEC 0" can, only talk to instruments where the “SEC nn” is a number ≥ 1.  The "SEC  nn" instruments can never be talked to directly from the controller.

Using a controller, an instrument can be returned to its native language (Compatible Language) by sending the string:

"SYST:LANG<space>COMP"

Without a controller, instruments which have been instructed to use SCPI can be changed back to the Compatible Mode using the following process:

Set rear address switch to "31".

Cycle the ac power, the display will show "Err  8".

Reset the rear address switch to "5".

Cycle the ac power, the display will show "Adr  5".

The instrument is now in its native language (Compatible language).

NOTES:

The choice of programming language is stored in the instrument’s nonvolatile ROM and remains the language of choice until purposefully changed by the user.

Errors reported as "8" of "-113" are characteristic of using a programming syntax which is inconsistent with the product’s settings/expectations.